## Slide 1: 參考電極注意事項 Notes for reference electrodes

## Slide 2: 量測參考電極(AgCl)注意事項 Precautions for Measuring Reference Electrode (AgCl)
使用參考電極前請先注意是否有氣泡，如有氣泡，請輕輕彈開 Please check whether there are bubbles in the reference electrode before using. If there are bubbles, please tap them off slightly.
測兩參考電極相對電壓，cell槽裡放保存液(3 M KCl)，以Open circuit Potential-Time測試系統檢查，假如相對電壓大於20mV，則其中一隻參考電極已受損 Measure the relative voltage of the two reference electrodes, put a preservation solution (3 M KCl) in the cell, and check with the Open circuit Potential-Time test system. If the relative voltage is greater than 20mV, one of the reference electrodes has been damaged
若是參考電極發現有異常，請以 0.5 M 氯化鉀溶液(KCl) 配置5mM 赤血鹽溶液 [K4Fe(CN)6]，以Cyclic Voltammetry測其標準電位 If the reference electrode is found to be abnormal, prepare a 0.5 M potassium chloride solution (KCl) with 5mM potassium ferrocyanide solution [K4Fe(CN)6], and measure its standard potential with cyclic voltammetry

## Slide 3: 量測參考電極(HgO)注意事項 Precautions for Measuring Reference Electrode (HgO)
使用參考電極前請先注意是否有氣泡，如有氣泡，請輕輕彈開 Please check whether there are bubbles before using the reference electrode. If there are bubbles, please bounce them slightly
測兩參考電極相對電壓，cell槽裡放保存液(1 M KOH)，以Open circuit Potential-Time測試系統檢查，假如相對電壓大於20mV，則其中一隻參考電極已受損 Measure the relative voltage of the two reference electrodes, put a storage solution (1 M KOH) in the cell, and check with the Open circuit Potential-Time test system. If the relative voltage is greater than 20mV, one of the reference electrodes has been damaged
若是參考電極發現有異常，請以 0.5 M 氯化鉀溶液(KCl) 配置5mM 赤血鹽溶液 [K4Fe(CN)6]，以Cyclic Voltammetry測其標準電位 If the reference electrode is found to be abnormal, please use 0.5 M potassium chloride solution (KCl) to prepare 5mM potassium ferrocyanate[K4Fe(CN)6], and measure its standard potential with Cyclic Voltammetry

## Slide 4: Cell 裝置
Open circuit Potential-Time
reference electrode
counter electrode
working electrode
reference electrode
counter electrode
working electrode
Cyclic Voltammetry
3 M KCl,
1 M KOH is OK.
5mM [K4Fe(CN)6]

## Slide 5: Open circuit Potential-Time

## Slide 6: Cyclic Voltammetry